## Slide 1: How to install external sequence at Siemens systems
Based on host IDEA version VE11
And FWF patch version 1.13 (and later)
Filip Szczepankiewicz, 2018-04-30

## Slide 2: Copy the sequence files to %customerseq%
Log in as advanced user. Normally by opening the start (press windows button, or Ctrl + Esc) and selecting ”Advanced User”
  Password is provided by Siemens
Insert the media that contains the sequence files (USB stick). Copy the relevant files to the customer sequence folder
  Usually, two files need to be copied (.dll and .so)
  Acess the correct folder by typing %customerseq% in an explorer window and press enter.
Hint: It is likely that %customerseq% points to C:/Medcom/MriCustomer/seq/

## Slide 3: Create a valid waveform library (not applicable if you have ”s” version)
Create a common library folde where generic waveforms are stored
  The path is: %customerseq%/FWF_Library/COMMON/
  This folder should be populated by
    FWF_OPTICUBE03.txt, FWF_OPTICUBE12.txt, FWF_OPTICUBE21.txt
    For > V1.07 it also includes FWF_TRAPXXX.txt and FWF_BITRAPXXX.txt
Insert your own waveforms in one of the study folders
  The path is: %customerseq%/FWF_Library/STUDY_*/
    The ”*” can be 01, 02,… 20
  In each study folder you can keep 16 pairs of waveform files
    These are named: FWF_CUSTOM*_A.txt and FWF_CUSTOM*_B.txt
    The ”*” can be be 001, 002, … 010
    The _A and _B are played before and after the refocusing pulse
Hint: Don’t modify waveform files that you do not own – it may ruin someone’s setup.

## Slide 4: Open the parameter editor
Click the little white arrow next to the dot cocpit icon
  The program editor icon should appear
Click theprogram editor icon
Tip: Sometimes the Dot Cocpit and the Program editor icons are switched. In that case, you don’t need to click the white arrow to expand the menu.

## Slide 5: Select a work folder
Navigate to an appropriate folder, and select it by clicking its name
  We call this the ”work folder”
Click the ”Edit” button to enter the program editor
Tip: To insert sequences, you must be at the lowest level of the system tree. There are three levels. In the example these are: WORKSHOP MDM/Install Demo/Group1

## Slide 6: Import the sequence from customer seq
Click the ”Browse” tab
Select ”Default” in drop down
In the list below, navigate to
  Sequence Region
  Customer Sequences
  Default
The sequence appears in the list.
Drag-and-drop the appropriate sequence into the work folder
  Only one object is moved this way, but both .dll and .so are installed by doing so.
Tip: If the sequence does not appear in the list, it may be compiled for the wrong software version.

## Slide 7: The sequence is now imported!
The imported exam card can now be used on the scanner
Note that the standard settings are not meant to be used for actual imaging
If you could not find the sequence in the previous step, it is possible that the sequence is incompatible with your system.